## Tabelle1
Einbauplatz | 3 | 4 | 5 | 6 | 7
Seriennummer | 22797839 | 22797840 | 22797837 | 22797836 | 22797838
PcbId | 222830038 | 222830333 | 222830037 | 222830371 | 222830336
Status | Messung beendet | Messung beendet | Messung beendet | Messung beendet | Messung beendet
Fehler [%] | 26216.57 | -0.29 | -0.85 | -0.45 | -0.88
Druchfluss [m³/h] | -14899.7356 | 56.4524 | 56.1357 | 56.3617 | 56.1182
Volumen [m³] | 329.9095 | 1.25 | 1.243 | 1.247 | 1.2396
Zeit [S] | 79.71 | 79.71 | 79.71 | 79.65 | 79.52
Fehlergrenze [%]
Zielwert Justage [%] | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
Pfad 1 Fehler [%] | -1.61 | 0.4 | -0.79 | -0.32 | -1.38
Pfad 1 Druchfluss [m³/h] | 55.7033 | 56.8426 | 56.1691 | 56.4368 | 55.836
Pfad 1 Volumen [m³] | 1.2334 | 1.2576 | 1.2437 | 1.2486 | 1.2344
Pfad 1 Zeit [S] | 79.71 | 79.65 | 79.71 | 79.65 | 79.59
Pfad 2 Fehler [%] | 0.51 | -1.3 | -1.41 | -2.61 | -2.37
Pfad 2 Druchfluss [m³/h] | 56.9062 | 55.8821 | 55.8183 | 55.1377 | 55.2727
Pfad 2 Volumen [m³] | 1.259 | 1.2364 | 1.235 | 1.2199 | 1.2219
Pfad 2 Zeit [S] | 79.65 | 79.65 | 79.65 | 79.65 | 79.59
Pfad 3 Fehler [%] | -0.04 | 0.03 | -0.35 | 1.46 | 0.21
Pfad 3 Druchfluss [m³/h] | 56.5961 | 56.6329 | 56.4191 | 57.4431 | 56.7366
Pfad 3 Volumen [m³] | 1.2531 | 1.253 | 1.2483 | 1.2709 | 1.2533
Pfad 3 Zeit [S] | 79.71 | 79.65 | 79.65 | 79.65 | 79.52

## Tabelle2
Cordonel Calculation Of Scale Factor
Date: | 2023-01-18
File: | CordonelCalFactorCalculationCheck.xlsx
Author: | Roland Drabesch/Thomas Wiedebusch
Nominal Scale Cordonel: | 15625 | UInt16
Required Deviation: | 1 | %
| Input of real measurements |  | Constant flow required for this |  | Measured deviation after Qm correction | [5 empty] | Back calculation using the real calculation factor
Name | Measurement Time | Qm | FlowRateCmPh | Qm Corrected | ScaleFactorRefToDut | DeviationDutToRefRel | DeviationDutToRefPer | reqDeviationRel | CalculateCalibFactorRel | BuildMeterCalibFactor | Qm corr | Dut to Ref Deviation Theoretical
Unit | s | m³ | m³/h | m³ | norm | norm | % | norm | norm | UInt16 | m³ | norm | %
Display | 79.6486 | 1.2428 | 56.1737 | 1.2483 | 1.0075 | -0.0074506 | -0.74506 | 1.01 | 1.0176 | 15900 | 1.2702 | 0.010018 | 1.0018 | 1
Channel 1 | 79.6486 | 1.2436 | 56.21 | 1.2491 | 1.0069 | -0.0068089 | -0.68089 | 1.01 | 1.0169 | 15889 | 1.2702 | 0.009972 | 0.9972
Channel 2 | 79.6486 | 1.2361 | 55.8702 | 1.2415 | 1.013 | -0.012812 | -1.2812 | 1.01 | 1.0231 | 15986 | 1.2702 | 0.009996 | 0.9996
Channel 3 | 79.711 | 1.2497 | 56.4398 | 1.2542 | 1.0028 | -0.0027482 | -0.27482 | 1.01 | 1.0128 | 15825 | 1.2702 | 0.010017 | 1.0017
Ref | 79.9973 | 1.2576 | 56.5953 | 1.2576 | 1 | 0 | 0
Qm Corrected using the flow rate of the device and the refrence time to calculate the accumulated volume back to the ref time
For theoretical back-calculation the Qm corrected will be taken.
The calculated calib factor corrects the deviation of the corrected volume and the required deviation
Base of the correction to the real measuremnt Om corrected is the requirement of a constant flow rate during the measurement time
Before each calibration run the default calib factor has to be set(15625)