|  | Software Requirement Specification | SENSUS GmbH |
| --- | --- | --- |
| Date: 17-Jul-2018 | Test Bench Framework (TBF) | Revision: A |
| History: Rev; Date; Comment; Author; Approved (Date/Signature). A; 07-Mar-2017 17-Jul-2018; Initial; T.Wiedebusch. |
|  | Copyright © 2017..2018 by SENSUS GmbH |  |
| File: TBF_SRSe.docx |  |  |

List of Contents

1	General Part	6

1.1	Introduction	6

1.2	Purpose	6

1.3	Content	6

1.4	System Description	6

1.5	System Overview	6

2	General Requirements	7

2.1	System Requirements	7

2.2	Development Tools	7

2.3	Coding Convention	7

2.4	Standard Naming	7

2.5	Units	7

2.6	Technical Guiding Rules	7

2.7	Documentation Language	8

2.8	Required Documents	8

3	TBF Requirements	9

3.1	Measurements	9

3.1.1	Static Start/ Stop Measurement	10

3.1.2	Flying Start/ Stop Measurement	11

3.2	Calibration Process	13

3.2.1	Preparation	13

3.2.2	Calculation of Calibration Factor	13

3.2.3	Finalization	13

3.3	Accuracy Validation Process	13

3.3.1	Preparation	13

3.3.2	Calculation of Accuracy	13

3.3.3	Finalization	14

3.4	Communication Chains	14

3.4.1	Request via IrDA	14

3.4.2	Request via RFID	14

3.4.3	Request via UART	14

3.4.4	Streaming the LED	14

3.5	Communication Protocols	14

3.5.1	Request Protocol	14

3.5.2	Streaming Protocol	15

3.5.2.1	Streaming Protocol Version ‘f’	15

3.5.2.2	Streaming Protocol Version ‘h’	16

3.6	Transmission Protocols	17

3.6.1	IrDA Transmit Protocol	17

3.6.2	RFID Transmit Protocol	18

3.6.3	UART Transmit Protocol	18

3.6.4	LED Transmit Protocol	18

3.7	Communication Ports	18

3.7.1.1	Serial Communication Port	18

3.7.1.2	IrDA Communication Port	18

3.7.1.3	RFID Communication Port	18

3.8	Record Types	18

3.8.1	ASCII String Record Type	18

3.8.2	Byte Record Type	19

3.9	Overflow Handling	19

3.9.1.1	Timing Overflow	19

3.9.1.2	Calibration Volume Overflow	19

3.9.1.3	Display Volume Overflow	19

4	Schedule	20

4.1	TBF Meeting Minutes	20

5	Contacts	20

6	References	21

List of Tables

Table 1.1 Abbreviations	5

Table 1.2 Addresses	5

Table 3.1 Streaming Protocol Version ‘f’	15

Table 3.2 Streaming Protocol Version ‘h’	17

Table 5.1 Contacts	20

List of Pictures

Picture 1.1 Block Sheet	6

Picture 3.1 Measurement Block Sheet	9

Picture 3.2 Flying Start/Stop Measurement Sequence	10

Picture 3.3 Flying Start/Stop Measurement Start-Sequence	11

Picture 3.4 Flying Start/Stop Stop-Sequence	12

Picture 3.4 Request via IrDA	14

Picture 3.4 Request via RFID	14

Picture 3.4 Request via UART	14

Picture 3.4 Streaming the LED	14

List of Equations

Equation 3.1 Average Flow Rate	12

Equation 3.2 Recalculated Test Volume	12

Equation 3.3 Calculation of Calibration Factor	13

Equation 3.4 Calculation of Accuracy	13

Abbreviations

| cmd | command |
| --- | --- |
| DLL | Dynamic Link Library |
| DUT | Device under Test, the water meter being manufactured or tested |
| Elde | Special hardware interface for test bench control |
| Genesis | Water meter with ultra-sonic measurement |
| FM2014 | Frequency measurement system |
| FM85/FM85.P | Frequency measurement system, legacy |
| FW | Firmware |
| GUI | Graphical User Interface |
| HW | Hardware |
| MID | Magneto Inductive Device, reference water meter |
| Prüf2000.exe | Laatzen production test bench software written in VB |
| PTB | Physikalisch Technische Bundesanstalt, German norming authority |
| P20xx | Test bench identifiers |
| SDD | Software Design Document |
| SIS | Software Interface Specification |
| SPS | Speicherprogrammierbare Steuerung, Siemens™ device for HW-Interfaces |
| SRS | Software Requirement Specification |
| SW | Software |
| tbd | to be defined |
| TBF | Test Bench Framework, this software product controlling the test bench |
| VB | Visual Basic |

Table 1.1 Abbreviations

Addresses

SENSUS GmbH, Meineckestr. 10, 30880 Laatzen, Germany

SENSUS GmbH, Industriestraße 16, 67063 Ludwigshafen am Rhein, Germany

SENSUS Metering Systems, Dr. Alberta Schweitzera 194, 91601 Stará Turá, Slovakia

Sentec Ltd, 5 The Westbrook Centre, Milton Road, Cambridge, CB4 1YG, UK

HSW Thomas Wiedebusch, Ostermarsch 17c, 31191 Algermissen, Germany

Table 1.2 Addresses

# General Part

## Introduction

This SRS document describes the TBF (V2) requirements. The initial software has been developed by Milan Hanajik to support world-wide test benches containing an Elde controller.

A special adaption is being needed to adapt this framework to the specific infrastructure of the SENSUS-Laatzen test benches, which do use a FM2014 and a SPS to control the test bench and to calibrate and test the water-meters instead of the Elde controller.

A new test bench is going to be installed for the Genesis-water-meter. This test bench initially uses the Elde controller and shall support later on the SPS + FM2014 as replacement for Elde.

New to the TBF is the module supporting the Genesis. This has to be developed by SENSUS-Laatzen. For Genesis support of the Laatzen production test benches, a plugin DLL has to be developed for the existing VB program Prüf2000.exe which does currently control these test benches.

A later development will import the entire Laatzen production test bench HW to the TBF being able to exchange the legacy Prüf2000.exe and bringing a world-wide TBF standard into the field.

## Purpose

For summarization of all work needed to be done, this document is being used as common discussion base. It will avoid misunderstandings of the requirements. Cyclically it shall be reviewed by certain authorities or experts to approve the requirements.

## Content

## System Description

## System Overview

Picture 1.1 Block Sheet

# General Requirements

The basic requirements will cover all elements needed to build and run the software. It should be avoided to enforce the operation at legacy systems which are not state of the art. A common list of tools may be needed to simplify data and file exchange or rebuild the software on a different PC.

## System Requirements

PC with Intel™ I5 Core or higher

Operating system MS-Windows™ 7 professional and higher

Minimum 2 GB RAM

Minimum 100 GB SDD

Display resolution 1280x1024 dots and higher

## Development Tools

Development Environment:		MS-VisualStudio™ 2015/2017 Professional

Compiler:				MS C# 6.0

.Net Framework:			4.0	(for WinXP compatibility)

Documentation:			MS Word/Excel/Visio™

Code Revision (CVS):		MS-TFS™ 2012

Test-Projects .Net Framework:	4.5.2

TBF .Net Framework:			4.5.2

## Coding Convention

It is required to use the English language for all source code comments. A coding standard shall be used for program internals.

## Standard Naming

The DIN EN14154 Part 1 to 3 [DIN_EN14154] describes the test procedures for water-meters. This norm can be found in German und English language. This norm is the base for naming conventions regarding the the specific vocabulary used in concatenation with the water-meters, the test bench and metrology.

The menu item language (English/German) shall be selectable for the country the program is going to be used. The table [MTS_1_3] will show the translations between English and German.

## Units

All units shall be based on SI-Units. A unit list for standardized SI-Units has been published by the PTB institute [SI_UNITS].

## Technical Guiding Rules

Guiding rules for water-meter-test have been published by the German PTB institute and are required for all test executed in Germany [TR_W2] and [DIN_EN14154].

## Documentation Language

It is required to use the English language for every documentation and the user manual.

## Required Documents

Software Requirement Specification (SRS)

Software Design Document (SDD)

User Manual (UM)

# TBF Requirements

## Measurements

Two different types of measurements can be executed to calibrate the DUT or validate the accuracy, the “Static Start/ Stop Measurement” or the “Flying Start/ Stop Measurement”.

The measurement block sheet [Picture 3.1] will give a rough simplified overview of how a software and hardware components are connected to each other and the command flow between these modules.

Picture 3.1 Measurement Block Sheet

### Static Start/ Stop Measurement

For the “Static Start/ Stop Measurement” the reference volume and the desired deviation is being needed from the overlying test program. The reference volume will directly read from a weight scale or the MID. This procedure is independent of any measurement time. The measured volumes of the reference and the DUT will directly be used to calculate the scale, as described in [Chapter 3.2.2 Calculation of Calibration Factor], or the DUT accuracy, as described in [Chapter 3.3.2 Calculation of Accuracy].

Picture 3.2 Flying Start/Stop Measurement Sequence

The measurement sequence will be:

- Send “Start cmd” to DUT batch,

- The DUT batch will prepare all connected DUTs,

- Wait for readiness of DUTs, these will send “Cyclic data” if ready,

- Read reference start volume from weight scale or assume 0 counted pulses for the MID,

- Send the “Start event” to the protocol decoding unit,

- “Save start values” of each DUT,

- Start the pump,

- “Get weight”  or “Get pulses” from the FM2014 until required volume has been collected,

- Stop pump and wait a certain time for zero-flow stabilization,

- Read the “Get pulses” from the FM2014 or take the volume of the weight scale,

- Send the “Stop event” to the protocol decoding unit,

- “Save stop values” of each DUT,

- Send “Stop cmd” to DUT batch, this will prepare all DUTs,

- “Calculation” of deviation between MID and DUT.

### Flying Start/ Stop Measurement

The required base of a “Flying Start/ Stop Measurement” is the constant flow rate during the entire measurement! This is essential for the enclosed procedure and the calculations of deviations and scale factors. The flow rate is going to be controlled by the SPS. Before the measurement can be started, the desired flow rate has to be reached and stabilized!

Picture 3.3 Flying Start/Stop Measurement Start-Sequence

The “Flying Start/Stop Measurement” start-sequence will be:

- Start the pump and wait for stabilization of the required flow-rate,

- Send “Start cmd” to DUT batch,

- The DUT batch will prepare all connected DUTs,

- Wait for readiness of DUTs, these will send “Cyclic data” if ready,

- Program the required pulses for the measurement volume to the FM2014 for the MID,

- Send the “Start event” to the protocol decoding unit,

- “Save start values” of each DUT.

- Measurement ongoing……

Picture 3.4 Flying Start/Stop Stop-Sequence

The “Flying Start/Stop” start-sequence will be:

- Wait until the remaining pulses from the FM2014 reached zero indicating the volume has been collected,

- Send the “Stop event” to the protocol decoding unit,

- “Save stop values” of each DUT,

- Send “Stop cmd” to DUT batch,

- The DUT batch will prepare all connected DUTs,

- Stop pump,

- Get the “reference time” for the exactly measurement duration from the FM2014,

- “Calculation” of deviation between MID and DUT.

For the “Flying Start/ Stop Measurement” the individual start- and stop-volume and the start- and stop-time is being needed from the overlying test program.

Out of the start and stop values of the DUT the flow rate will be calculated.

QDUT= VDUTstop-VDUTstarttDUTstop-tDUTstart

Equation 3.1 Average Flow Rate

The reference time is being used to recalculate the volume during the test time, due to the fact that the start and stop data record may be delayed by the software program flow or the unpredictable operational system timing.

VDUT= QDUT tREF

Equation 3.2 Recalculated Test Volume

## Calibration Process

The Genesis water meter has to be calibrated to get accurate readings. In advance the amplitude and zero-flow-offset correction has to be done.

Due to the 3 paths of the Genesis, these have to calibrated independently. Each of these paths measures the entire volume as individual circuitry. The result is the average of all three paths.

### Preparation

TODO Display volume

Calibration scale (default)

Switch LED to calibration mode

### Calculation of Calibration Factor

The calibration is being done using the measured volume of each individual path, the reference volume, measured by a weight scale or a MID, and for the flying start/stop measurement the timings of the individual channel and the precisely timing of the reference system.

These values will be used to calculate the new calibration factor sx. It is based on the default scale factor  s_defx multiplied by the desired deviation k at this specific flow-rate. The desired deviation is needed to adjust the curve of deviations depending on different flow rates into an envelope of tolerances. Detailed information of the calibration procedure for the Genesis-water-meter can be found in [GenCal].

sx= VREFVDUTx s_defx  k          (x=1..3;  0.95<k<1.05)

Equation 3.3 Calculation of Calibration Factor

### Finalization

TODO check tolerance

Switch LED off

## Accuracy Validation Process

To verify the function of the device for different flow rates, several validation measurements have to be executed.

### Preparation

TODO Display volume

Switch LED to calibration mode

### Calculation of Accuracy

The deviation dQ at a specific flow rate shall be calculated using the reference and the measured volume.

dQ= VrefVx

Equation 3.4 Calculation of Accuracy

### Finalization

TODO check tolerance

Switch LED off

## Communication Chains

The following chapter shows the different layers of the SW for the communication of different protocols and HW-interfaces.

### Request via IrDA

Picture 3.5 Request via IrDA

### Request via RFID

Picture 3.6 Request via RFID

### Request via UART

Picture 3.7 Request via UART

### Streaming the LED

Picture 3.8 Streaming the LED

## Communication Protocols

The communication protocols describe the elementary interface to the Genesis independend of the underlaying transmittion protocol or HW-Interface.

### Request Protocol

The “Request Protocol” has to be used for the interactive communication to the Genesis. This protocol supports the login procedures, read out of information, change of register values for setup, test or calibration purposes. The complete descripton can be found in the “Genesis Meter Configuration.docx” [GenConf].

Example for “Read Parameter” of Genesis water meter calibration factor1:

Request

| cmd | register |  | data |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 0x09 | 0x0D | 0x0F | 0x00 | 0x00 | 0x00 | 0x00 |

Answer

| cmd | error code |  | data |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 0x0A | 0x00 | 0x00 | 0x34 | 0x12 | 0x00 | 0x00 |

### Streaming Protocol

The “Streaming Protocol” is being used to receive data from the Genesis. These can be switched to different output versions. These output versions are described in the “Genesis Debug Serial Output.dotx” [GenDbg]. Each output version contains a set of data needed for test, debug or calibration purposes.

#### Streaming Protocol Version ‘f’

This protocol will be used for the “Accuracy Validation Procedure”.

| Field Number | Description | Details | Unit / Scale | Resolution | Validated |
| --- | --- | --- | --- | --- | --- |
| 1 | Display volume | The volume as displayed on the LCD. Note this will be in the units as displayed on the LCD. | depending on setup ml, gallons, m³ |  | x |
| 2 | Time | A timestamp in seconds scaled up by 16 bits (0x10000 = 1s) | µs/digit | 15.25879 | x |
| 3 | CRC | The CRC of all of the characters (separators included). The CRC is the 16bit CCITT with polynom 0x1021 MSB first, initialised with 0xFFFF |  | x | x |

Table 3.1 Streaming Protocol Version ‘f’

#### Streaming Protocol Version ‘h’

This protocol will be used for the “Calibration Procedure” and raw value logging during “Accuracy Validation Procedure”.

| Field Number | Description | Details | Unit / Scale | Resolution(R) Valid range(V) Min/Max(M) Error(E) Example(X) | Vali-dated |
| --- | --- | --- | --- | --- | --- |
| 1 | Channel number | 1, 2 or 3 | x |  |  |
| 2 | Error | An error code describing the following data 0 = No error 0x0F0B = Temperature below range 0x0F0C = Temperature above range 0x0F0D = Bad pulse width difference 0x0F0E = Amplitude out of range 0x0F0F = Time of flight out of range 0x0F10 = Delta time of flight out of range 0x0F11 = Time of flight measurement timeout 0x0F12 = Time of flight out of range |  |  |  |
| 3 | Total time of flight | The average delta time of flight for the current measurement represented in internal units scaled by 38 bits (0x40 0000 0000) | ps/digit | R:3.63798 |  |
| 4 | Delta time of flight | The average delta time of flight for the current measurement represented in internal units scaled by 38 bits (0x40 0000 0000) | ps/digit | R: 3.63798 |  |
| 5 | Raw delta volume | The volume calculated in the last interval on this channel | digit |  |  |
| 6 | Raw accumulated volume | The accumulated volume for this channel | digit |  |  |
| 7 | Raw volume scaling | The value of raw volume that equates to 1ml | digit/ml | X:1024 |  |
| 8 | Sample interval | The amount of time since the last sample in seconds scaled up by 16 bits (0x10000 = 1s) | µs/digit | R:15.25879 |  |
| 9 | Amplitude up | The amplitude of the receive signal in mV scaled up by 22 bits (0x40 0000 = 1mV) | pV/digit | R:238.419 V:200…650mV |  |
| 10 | Amplitude down | The amplitude of the receive signal in mV scaled up by 22 bits (0x40 0000 = 1mV) | pV/digit | 238.419 V:200…650mV |  |
| 11 | Temperature | The temperature in degrees C scaled by Temperature scale below | digits/ 2^TempScale |  |  |
| 12 | Temperature scale | Number of fractional bits in Temperature | power factor |  |  |
| 13 | Time | A timestamp in seconds scaled up by 16 bits (0x10000 = 1s) | µs/digit | 15.25879 |  |
| 14 | CRC | The CRC of all of the characters (separators included). The CRC is the 16bit CCITT with polynom 0x1021 MSB first, initialised with 0xFFFF |  |  |  |

Table 3.2 Streaming Protocol Version ‘h’

## Transmission Protocols

The transmission protocols are strongly dependend of the underlaying HW-Interface. Each interface has a specific protocol at the transmission layer containing as payload the “Request Protocol” or “Streaming Protocol”.

### IrDA Transmit Protocol

The detailed definition of the IrDA protocol can be found in [IrDADef].

This protocol will be used for the “Request Protocol”. It uses the “Byte Record Type” for data transmission. The “Request Protocol” can be found in the payload range of the “IrDA Transmit Protocol”.

Example for “Read Parameter” of Genesis water meter calibration factor1:

Request

| IrDA header |  |  |  | IrDA payload, Request Protocol |  |  |  |  |  |  | IrDA CRC |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| sync | head | length | ocmd | cmd | register |  | data |  |  |  |  |  |
| 0x9B | 0x20 | 0x07 | 0x02 | 0x09 | 0x0D | 0x0F | 0x00 | 0x00 | 0x00 | 0x00 | 0xXX | 0xXX |

Answer

| IrDA header |  |  |  | IrDA payload, Request Protocol |  |  |  |  |  |  | IrDA CRC |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| sync | head | length | omsg | cmd | error code |  | data |  |  |  |  |  |
| 0x9B | 0x61 | 0x07 | 0x03 | 0x0A | 0x00 | 0x00 | 0x34 | 0x12 | 0x00 | 0x00 | 0xXX | 0xXX |

sync:	synchronization byte of IrDA, always 0x9B

head: 	header information of frame type, direction and listen after talk

length: length of the payload

ocmd:	optical command for request, always 0x02

omsg:	optical message for answer, always 0x03

### RFID Transmit Protocol

This protocol is obsolete.

The “RFID Transmit Protocol” will be used for the “Request Protocol”. It uses the “Byte Record Type” for data transmission.

The CRC calculation does leave the CRC field out.

### UART Transmit Protocol

This protocol is obsolete.

The “UART Transmit Protocol” will be used for the “Request Protocol”. It uses the “Byte Record Type” for data transmission.

The CRC calculation includes the CRC field itself filled with 0’s in advance.

### LED Transmit Protocol

This protocol will be used for the “Streaming Protocol”. It uses the “ASCII String Record Type” for transmission of data.

## Communication Ports

The Genesis water-meter uses two communication ports for test and calibaration purposes. Depending in the revision and HW these can be directly accessed, mentioned as “Serial Communication Port”, or redirected to a HW-Interface like “IrDA Communication Port”.

#### Serial Communication Port

This communication port is obsolete for the “Request Protocol” but will be used for the “Streaming Protocol” in concatination with the “LED Transmit Protocol”.

#### IrDA Communication Port

The detailed definition of the IrDA protocol can be found in [IrDADef].

An additional wake-up burst two times 0xF0 with subsequent delay time of at least 500µs before dispatching the “IrDA Transmit Protocol” record.

#### RFID Communication Port

This communication port is obsolete.

Record split to message chunks of 6 bytes RFID payload.

## Record Types

Two different record type formats are going to be used by the Genesis water meter.

### ASCII String Record Type

String transmission indicated by a line end delimiter “carriage return”. This record type will be used for the “Streaming Protocol”.

### Byte Record Type

This record type will be used for the “Request Protocol”. It contains at least a synchronization byte as first byte of the record and a record length.

## Overflow Handling

Due to the limited size of the values to 32Bit integer or unsigned integer values, an overflow has to be taken into account and carefully handled by the software.

#### Timing Overflow

The timing resolution is given by 1s / 2^16digits = 15.25879µs/digit. Due to the value format of 32Bits unsigned integer, an overflow will take palce after 65,536s or 18h:12min:16s.

#### Calibration Volume Overflow

The resolution for the 3 paths raw value is based on a 32Bit unsigned integer and a resolution of 1ml/1024digits. This results in an overflow volume of  4.194304m³.

#### Display Volume Overflow

After the display of 999.999 999m³ at a setup of 1ml resolution this value will overflow to 0.0m³ (keeping ineternally the overflow into account).

# Schedule

## TBF Meeting Minutes

The link to the meeting minutes will show the actual work to be done.

http://project.emea.sensus.net/PWA/EMEAAP_Genesis/Project Documents/300_R+D/380_Product_Test+Benchmarks/384_Genesis_Testbench/MeetingMinutes_GenesisTbfFreitagsRunde.xlsm

# Contacts

All mail-addresses can be build using the firstName.lastName@xyleminc.com unless otherwise noted.

| Name | Function | Tel / mail |
| --- | --- | --- |
| Albert Morwig |  | +49 621 6904 1286 +49 171 2268 958 |
| Thomas Helf |  |  |
| Thomas Quedenbaum | Leader production | +49 5102 743 175 |
| Peter Buch |  | +49 5102 743 183 |
| Jens Lippold |  | +49 5102 743 180 |
| Thomas Wiedebusch | Software Developer | +49 179 146 04 93 thomas.wiedebusch@t-wiedebusch.de |
| Roland Drabesch | Software Developer | +49 5102 743 225 |
| Reinhard Henning | Software Developer | +49 171 2859 742 |
| Milan Hanajik | Primary software designer for TBF | +42 1904 195 145 milan.hanajik@hotmail.com |
| Nikolas Orelly | IT |  |
| Karsten Nettemann | Test bench expert/ Beta tester | +49 5102 743 150 |
| Peter Sandvoß | Test bench expert/ Hardware specialist | +49 5102 743 118 |

Table 5.1 Contacts

# References

[SI_UNITS]	SI-Units, Published:  Physikalisch Technische Bundesanstalt (PTB),  117.Edition, Booklet 2, June 2007

[TR_W2]	TR W2 Technische Richtlinien für Wasserzähler, Published:  Physikalisch Technische Bundesanstalt (PTB),  Edition December 2012

[DIN_EN14154]	DIN EN14154 Water meters test methods and equipment, Published: DIN Deutsches Institut für Normung e.V., Part 3, Edition July 2007

[MTS_1_3]	Translation MTS 1_3, Published: SENSUS GmbH, Edition 16. March 2011

[GenDbg]	Genesis Debug Serial Output; Published:  Sentec; Author: Alex Frost

http://project.emea.sensus.net/PWA/EMEAAP_Genesis/Project%20Documents/300_R%2BD/340_Software%2BFirmware/342_Software_and_Firmware/Genesis%20Debug%20Serial%20Output.docx

[GenCal]	Calibration Procedure; Published:  Sentec; Authors: Alex Frost, Ben Davey, Joe Adam

http://project.emea.sensus.net/PWA/EMEAAP_Genesis/Project%20Documents/300_R+D/340_Software+Firmware/342_Software_and_Firmware/Metrology%20and%20calibration%20firmware/Calibration%20procedure.docx

[GenConf]	Genesis Meter Configuration; Published:  Sentec; Author: Alex Frost

http://project.emea.sensus.net/PWA/EMEAAP_Genesis/Project%20Documents/300_R%2BD/350_Electronics_Radio/356_production%20processes/Genesis%20Meter%20Configuration.docx

[IrDADef]	NA2W IrDA Protocol Specification; Published:  SENSUS; Author: Kevin Shay, Adam Hansen

https://intranet.sensus.net/drc/engineering/water-gas/FW/Shared Documents